VCCS Command Pilot™ Inline — Engineering Brief

Prospective project overview for engineering partners     |     Version 1.0  ·  3 August 2026     |     DRAFT — subject to revision

1.  What is VCCS Command Pilot™?

Command Pilot is a Windows desktop application — built on WPF / .NET 8 — that VCCS/PIPS operates in its verification laboratory to assess barcode print quality for pharmaceutical, medical device, and other customers.  It accepts scan results from a Cognex DataMan verifier, decodes GS1 linear and 2D payloads (and direct-part marks), grades them against ISO/IEC 15415 and related standards, logs each result to Excel and an event log, and presents the outcome to a lab operator.  The application runs on a bench-top verifier station; it currently runs on a bench-top verifier station, not on a production line.

Command Pilot Inline is the inline production variant.  The same proven engine handles decode and grading; a new operator-facing panel and a new I/O relay assembly are added to close the loop with the physical line.

2.  Platform architecture (relevant assemblies)

DeviceInterface — Cognex DataMan SDK integration (port 44444), raw DMCC command client (port 23), trigger management, image acquisition.  Handles all reader communication; the rest of the stack never touches hardware directly.

ExcelEngine — Structured verification-record writer.  Produces session Excel workbooks (.xlsx / .xls) and an append-only JSONL event log.  Schema is configurable; this writer would likely be used for inline operation, though other reporting options are also possible.

ConfigEngine — Operator-configurable parameters managed at multiple security levels. Routine job settings (job name, operator ID, roll number, etc.) are accessible to production operators; grading thresholds and other quality-critical parameters are controlled at a more secure level, separate from day-to-day job configuration.

InlineIo  (new) — Relay-board abstraction layer.  Exposes IRelayBoard (hardware-agnostic), MockRelayBoard (development/test), IndicatorPoleController, and ConveyorInterruptController.  Designed to accommodate indicator pole, pusher-divert, emergency-stop (e-stop), and other output channels as assignments are confirmed by Engineering.  Exact line-control behaviour TBD.

Command Pilot Inline Operator Panel  (new) — Full-screen WPF panel designed for production-floor use: grade result, indicator status, session counters, operator controls.  Layout TBD with customer validation.

3.  How this applies to the inline application

The physical target is small folding cartons — roughly cell-phone to paperback-book footprint, 1.25–1.75 in. thick — conveyed flat with the narrow edge leading.  These are virus and other diagnostic test kits.  A Cognex DM-475V-LBL (adapted from the customer's existing desktop stand) reads the GS1 DataMatrix label as each carton passes.

The decode-and-grade loop is unchanged from the existing platform:

Photosensor (or alternative trigger — TBD) signals carton presence.

DeviceInterface fires a software or hardware trigger to the reader.

Reader decodes and returns the GS1 DataMatrix result + ISO/IEC 15415 grade.

ExcelEngine appends a verification record; image archived if required.

Grade result is handed to InlineIo for indicator and line-control output.

Indicator pole colour and line-control response are driven by the grade result.  The current working assumption is a multi-colour indicator pole and at least one pusher-divert output for failing cartons; whether a failing scan stops the line, diverts the carton while the line continues, or triggers some other response is a line-control integration question to be resolved with engineering.  The InlineIo layer is deliberately hardware-agnostic so these decisions do not require code changes.

4.  Reuse and new-build scope

| Component | Status | Confidence |
| --- | --- | --- |
| DeviceInterface (reader comms, trigger, image) | Reuse as-is | High |
| ExcelEngine (Excel + JSONL reporting) | Reuse as-is | High |
| ConfigEngine (operator config, PIN lock) | Reuse as-is | High |
| InlineIo relay board layer | New — stub complete | — |
| Command Pilot Inline Operator Panel (WPF) | New build | — |
| Line-control integration (divert, stop, etc.) | TBD Engineering | — |
| Relay board hardware + wiring | TBD Engineering | — |

5.  Key open items for engineering

Relay board: Model, channel count, USB/serial interface — drives IRelayBoard implementation.

Pusher divert: Output type (relay, PLC digital out), pulse vs. latched, timing relative to carton position.  At least one divert channel assumed; exact behaviour TBD.

Line-control mode: Fail = divert only (line keeps running) vs. fail = divert + stop.  Consecutive-fail threshold (if any) TBD.  Red indicator may be momentary rather than latched — confirm with customer.

Indicator pole: Lamp count, colour mapping, steady vs. flash behaviour per grade band, and audible alarm channel.  Current mapping is a working assumption; all thresholds negotiable.

Trigger: Photosensor type and placement, end-of-carton detection (second sensor or decode event), debounce timing.

Product recipe / config: A per-product configuration file (approx. 6 products) is needed, capturing fixed product data (GTIN), grade pass/fail threshold, and line settings specific to that product.  Threshold is not a global setting — it is configurable per product by authorised personnel.

Image policy: Which frames to archive (all / fail-only / none), at what resolution and compression quality, storage path and retention policy, and DMST upload workflow for failures.

Reader mounting: Stand adaptation for conveyor height and carton size; working distance and FOV verification against carton footprint.

This document describes the current working concept.  All line-control behaviours, channel assignments, and grade thresholds are provisional and subject to revision as engineering requirements are confirmed.  Contact: George Wright IV — VCCS / Product Identification and Processing Systems, Inc. — gw4@pips.com